## 1.Section 8
PROJEK:
SECTION 8: SCHEDULE OF TECHNICAL DATA OF EQUIPMENT OFFERED |  | SCHEDULE OF TECHNICAL DATA OF EQUIPMENT OFFERED
The tenderer must complete this Schedule of Technical Data as set out below. All information asked for must be given correctly and supported by copies of manufacturer's published technical data for the equipment being offered.
Tenderer
Company
Cost (RM)
Weeks of Completion
ITEMS |  |  |  |  | JKR REQUIREMENT | Contractor's Proposal
A. | DENTAL SPECIALIST'S UNIT
Manufacturer
Country of Origin
Model
| Registration under Medical Device Act 2012 (please provide the supporting document). |  |  | Yes/No | YES
1 | DENTAL CHAIR
a | Types of height adjustment. |  |  |  | Motorised / Hydraulic
b | Provide smooth with non- jerk start |  |  | Yes/No | YES
c | Adjustable height of the seat. (cm) up to of 200 kg distributed mass. |  |  |  | 44-70 cm
d | Seamless upholstery of durable. Please specify material used:
e | Two foot switches to control movements. |  |  | Yes/No | YES |  | Moto+G24rised / Hydraulic
f | Supplied with movable and/or detachable armrest for easy patient access and wheelchair transfer. |  |  | Yes/No | YES
g | Equipped with electric safety cut-off switch, in case of collision / obstruction with any object. |  |  | Yes/No | YES
h | Tilting movements are synchronised with inclination of backrest and legrest. |  |  | Yes/No | YES
i | Chair position programmed with 4 working positions. Please specify:
| (a) Exit and entry. |  |  | Yes/No | YES
| (b) Automatic rinsing position. |  |  | Yes/No | YES
| (c) Auto return to last working position. |  |  | Yes/No | YES
| (d) At least two working positions according to operator preference. |  |  | Yes/No | YES
j | All programmes controlled on both dentist’s and assistant’s side. into the seat. Yes/No |  |  | Yes/No | YES
k | Internal Umbilical Type into the seat. |  |  | Yes/No | YES
l | Internal Junction Box into the seat. Yes/No |  |  | Yes/No | YES
m | Electrical safety shall comply to MS IEC 60601-1 standards or equivalent (please provide the supporting document). |  |  | Yes/No | YES
n | Any other features:
2 | DENTAL UNIT
a | Type of Display |  |  |  | Digital
b | Two filters for air and water |  |  | Yes/No | YES
c | Foot Control for:
| i) Start and stop of turbine handpiece, micromotor, scaler and electrosurgery. |  |  | Yes/No | YES
| ii) Speed of turbine, micromotor, scaler and electrosurgery. |  |  | Yes/No | YES
| iii) Drilling with/without water |  |  | Yes/No | YES
| iv) For positioning of chair. |  |  | Yes/No | YES
2.1 | Assistant’s Instrumentation
| 2.1.1 | Control Switches
| a | Must be touch pad. |  | Yes/No | YES
| b | Easy cleaning & disinfection. |  | Yes/No | YES
| 2.1.2 | Multiple Function Syringe
| a | Function for: | Air | Yes/No | YES
|  |  | Water | Yes/No | YES
|  |  | Spray | Yes/No | YES
| b | Syringe nozzle autoclavable up to 135°C. |  | Yes/No | YES
| c | Syringe nozzle easily disconnected. |  | Yes/No | YES
| d | Syringe nozzle non-drip and non water retraction. |  | Yes/No | YES
| e | Degree of rotation for the Syringe Nozzle. (degree) |  |  | 360O
| f | Syringe tubing can disconnected. |  | Yes/No | YES
| g | Any other features:
| 2.1.3 | Saliva Ejector
| a | Tip autoclavable up to 135°C. |  | Yes/No | YES
| b | Degree of rotation for the Tip. (degree) |  |  | 360O
| c | Holder designed for easy disinfection/cleaning. |  | Yes/No | YES
| d | Supplied with universal adapter for disposable suction tip. |  | Yes/No | YES
| 2.1.4 | High Vacuum Suction
| a | Nozzle easily disconnected. |  | Yes/No | YES
| b | Nozzle autoclavable up to 135°C. |  | Yes/No | YES
| c | Degree of rotation for the Nozzle. (degree) |  |  | 360O
| d | Suction intensity can be controlled. |  | Yes/No | YES
| e | Supplied with 2 spare tips (Narrow & Wide nozzle) |  | Yes/No | YES
| f | Light barier switch or button switch for holder. |  | Yes/No | YES
| g | Suction tubing easily disconnected. |  | Yes/No | YES
| h | Continuous suction for not less than 2 hours. |  | Yes/No | YES
| i | Any other features:
2.2 | Doctor’s Instrumentation
| 2.2.1 | Instrumentation
| a | Carried by an adjustable swivel arm with braking system. |  | Yes/No | YES
| b | Connected to the dental chair base moving parallel to the chair. |  | Yes/No | YES
| c | Any other features:
| 2.2.2 | Control Switches
| a | Laminated / Foil Switch. |  | Yes/No | YES
| 2.2.3 | X-Ray Viewer
| a | Attached / Built In |  | Yes/No | YES
| 2.2.4 | Instruments Tubings
| a | Easy disconnected for disinfection. |  | Yes/No | YES
| 2.2.5 | High Torque Turbine
| a | Make
| b | Model
| c | Degree of rotation for the Handpiece. (degree) |  |  | 360O
| d | Handpiece autoclavable up to 135°C. |  | Yes/No | YES
| e | Handpiece non-drip and non-retraction. |  | Yes/No | YES
| f | Handpiece and tubing should be with fibre optic system. (Not less than 25,000 Lux) |  | Yes/No | YES
| g | Handpiece must be push button chuck system. |  | Yes/No | YES
| h | Speed of the handpiece (rpm) |  |  | ≥ 350,000 rpm
| i | Easy Cleaning. Please specify material used. |  | Yes/No | YES
| j | Supplied with three (3) handpieces. |  | Yes/No | YES
|  | (a) Two (2) high torque handpieces. |  | Yes/No | YES
|  | (b) One (1) high torque miniature head turbine handpiece. |  | Yes/No | YES
| k | Any other features:
| 2.2.6 | Micromotor
| a | Make
| b | Model
| c | Type of Micromotor |  |  | Electric
| d | Speed (min. 2,000 rpm: max. 40,000 rpm) |  |  | ≥ 2000 - 40000 rpm
| e | Micromotor should be with LED fibre optic system. |  | Yes/No | YES
| f | Maintain a high torque at all intensity |  | Yes/No | YES
| g | Motor cannot be demagnetised |  | Yes/No | YES
| h | Clockwise and anticlockwise rotation |  | Yes/No | YES
| i | Internal cooling water spray |  | Yes/No | YES
| j | Must be supplied with the following handpieces (3 units):
|  | (a) One (1) contra-angle handpieces 1 : 1 with light system. |  | Yes/No | YES
|  | (b) Two (2) straight handpiece 1 : 1. |  | Yes/No | YES
| k | All handpieces autoclavable up to 135°C. |  | Yes/No | YES
| l | All contra-angle handpiece fibre optic system |  | Yes/No | YES
| m | All handpiece integrated water spray. |  | Yes/No | YES
| n | Any other features:
| 2.2.7 | Light Cure Unit
| a | Make
| b | Model
| c | Light Switch located at the handpiece of the unit. |  | Yes/No | YES
| d | Automatic signal sounds can be programmed. |  | Yes/No | YES
| e | Light Intensity must be minimum 1,000 mW/cm2. |  |  | ≥ 1,000 mW/cm2
| f | Light source - LED. |  | Yes/No | YES
| g | Supplied with an optical filter. |  | Yes/No | YES
| 2.2.8 | Ultrasonic Scaler
| a | Make
| b | Model
| c | Magneto-Strictive System / Piezometer System with 28,000 Hz or more. |  | Yes/No | YES
| d | Handpiece or handpiece sleeves autoclavable up to 135°C. |  | Yes/No | YES
| e | 6 detachable working tips for scaling. |  | Yes/No | YES
|  | (a) Buccal / lingual and proximal ( 2 tips ). |  | Yes/No | YES
|  | (b) Sub-gingival ( 4 tips ). |  | Yes/No | YES
| f | Supply with additional one scaler handpiece. |  | Yes/No | YES
| g | To be equipped with light LED fibre optic . |  | Yes/No | YES
| h | Any other features:
| 2.2.9 | Intra Oral Camera
| a | Make
| b | Model
| c | Made of stainless steel or equivalent |  | Yes/No | YES
| d | Image sensor : High sensitivity CMOS or CCD. |  | Yes/No | YES
| e | Resolution (min 752x582 pixels) |  | Yes/No | YES
| f | High resolution /magnification (min 40x) |  | Yes/No | YES
| g | Non-drip and non-water retraction. |  | Yes/No | YES
| h | Material of handpiece: |  | Please state | To state by tenderer
| i | Multimedia System:
|  | i) Suitable Processor such as i3 / i5 Intel or Dual Core or better. |  |  | To state by tenderer
|  | ii) Microsoft Window that can operate the system. |  |  | To state by tenderer
|  | iii) Minimum 4GB DDRAM or higher. |  |  | To state by tenderer
|  | iv) Minimum 500GB HDD hard drive or higher. |  |  | To state by tenderer
|  | v) A flat screen monitor built on to the dental unit |  |  | To state by tenderer
|  | vi) Accessories like keypad, mouse and etc shall be included. |  |  | To state by tenderer
| j | Any other features:
| 2.2.10 | Multiple Function Syringe
| a | Make
| b | Model
| c | Function for: | Air | Yes/No | YES
|  |  | Water | Yes/No | YES
|  |  | Spray | Yes/No | YES
| d | Handpiece autoclavable up to 135°C. |  | Yes/No | YES
| e | Degree of rotation for the Handpiece. (degree) |  |  | 360O
| f | Nozzle easily disconnected. |  | Yes/No | YES
| g | Non-drip and non-water retraction. |  | Yes/No | YES
| h | Material of handpiece: |  | Please state | To state by tenderer
| i | Any other features:
2.3 | Cuspidor Unit
a | Attached to the left side of the chair and moves together with the chair. |  |  | Yes/No | YES
b | Consist of spittoon, automatic tumbler filler, saliva ejector and high volume suction. |  |  | Yes/No | YES
c | Electric safety cut-off/limit switch. |  |  | Yes/No | YES
d | Durable spittoon, easy cleaning and cross infection control. |  |  | Yes/No | YES
e | Spittoon bowl one piece type. |  |  | Yes/No | YES
f | Saliva ejector and high volume suction swivelled for easy access by the assistant. |  |  | Yes/No | YES
g | Automatic Control filter activated by foil switches for tumbler. |  |  | Yes/No | YES
h | Fitted with a waste trap. |  |  | Yes/No | YES
i | Any other features:
2.4 | Operating Light
a | Make
b | Model
c | Synchronised movement with chair in vertical direction. |  |  | Yes/No | YES
d | Spring loaded arm with locking brake and height adjustment. |  |  | Yes/No | YES
e | Type of reflector: |  |  |  | Parabolic mirror
f | Light intensity : (lux) |  |  |  | 9,000 - 24,000 lux
g | Halogen bulb or LED of low voltage and easy to replace. |  |  | Yes/No | YES
h | Colour temperature at all intensity : (Kelvin) |  |  |  | 5,500 K
i | Operated by foot control or sensor touch control switch. Please specify |  |  | Yes/No | YES
j | Any other features:
2.5 | Operating Stool
a | Stool for operation : Good lumbar support and adjustable backrest. |  |  | Yes/No | YES
b | Stool for assistant : Good lumbar support and adjustable backrest which can function as armrest and adjustable. |  |  | Yes/No | YES
c | Wide stable base on 5 castors. |  |  | Yes/No | YES
d | Height adjustable: (mm)
e | Any other features:
2.6 | Independent Water Storage (Self Contained Water System) |  |  |  | YES
a | Make
b | Size
c | Material made
2.7 | Electro Surgery
a | Make
b | Model
c | Utilised/applied to all type of electrosurgery procedures |  |  | Yes/No | YES
d | Fitted with removal plug in connection |  |  | Yes/No | YES
e | Supplied with a complete set of electrode |  |  | Yes/No | YES
f | Electrode cable utoclave up to 135⁰C |  |  | Yes/No | YES
g | High frequency generator 1.5 to 2.0 MHz and output 25W-50W |  |  | Yes/No | YES
h | Any other features:
B. | DENTAL OFFICER'S UNIT
Manufacturer
Country of Origin
Model
| Registration under Medical Device Act 2012 (please provide the supporting document). |  |  | Yes/No | YES
1 | DENTAL CHAIR
a | Types of height adjustment. |  |  |  | Motorised / Hydraulic
b | Provide smooth with non- jerk start |  |  | Yes/No | YES
c | Adjustable height of the seat. (cm) up to of 200 kg distributed mass. |  |  |  | 44-70 cm
d | Seamless upholstery of durable. Please specify material used:
e | Two foot switches to control movements. |  |  | Yes/No | YES
f | Supplied with movable and/or detachable armrest for easy patient access and wheelchair transfer. |  |  | Yes/No | YES
g | Equipped with electric safety cut-off switch, in case of collision / obstruction with any object. |  |  | Yes/No | YES
h | Tilting movements are synchronised with inclination of backrest and legrest. |  |  | Yes/No | YES
i | Chair position programmed with 4 working positions. Please specify:
| (a) Exit and entry. |  |  | Yes/No | YES
| (b) Automatic rinsing position. |  |  | Yes/No | YES
| (c) Auto return to last working position. |  |  | Yes/No | YES
| (d) At least two working positions according to operator preference. |  |  | Yes/No | YES
j | All programmes controlled on both dentist’s and assistant’s side. into the seat. Yes/No |  |  | Yes/No | YES
k | Internal Umbilical Type into the seat. |  |  | Yes/No | YES
l | Internal Junction Box into the seat. Yes/No |  |  | Yes/No | YES
m | Electrical safety shall comply to MS IEC 60601-1 standards or equivalent (please provide the supporting document). |  |  | Yes/No | YES
n | Any other features:
2 | DENTAL UNIT
a | Type of Display |  |  |  | digital
b | Two filters for air and water |  |  | Yes/No | YES
c | Foot Control for:
| i) Start and stop of turbine handpiece, micromotor, scaler and electrosurgery |  |  | Yes/No | YES
| ii) Speed of turbine, micromotor, scaler and electrosurgery |  |  | Yes/No | YES
| iii) Drilling with/without water |  |  | Yes/No | YES
| iv) For positioning of chair. |  |  | Yes/No | YES
2.1 | Assistant’s Instrumentation
| 2.1.1 | Control Switches
| a | Must be touch pad. |  | Yes/No | YES
| b | Easy cleaning & disinfection. |  | Yes/No | YES
| 2.1.2 | Multiple Function Syringe
| a | Function for: | Air | Yes/No | YES
|  |  | Water | Yes/No | YES
|  |  | Spray | Yes/No | YES
| b | Syringe nozzle autoclavable up to 135°C. |  | Yes/No | YES
| c | Syringe nozzle easily disconnected. |  | Yes/No | YES
| d | Syringe nozzle non-drip and non water retraction. |  | Yes/No | YES
| e | Degree of rotation for the Syringe Nozzle. (degree) |  |  | 360O
| f | Syringe tubing can disconnected. |  | Yes/No | YES
| g | Any other features:
| 2.1.3 | Saliva Ejector
| a | Tip autoclavable up to 135°C. |  | Yes/No | YES
| b | Degree of rotation for the Tip. (degree) |  |  | 360O
| c | Holder designed for easy disinfection/cleaning. |  | Yes/No | YES
| d | Supplied with universal adapter for disposable suction tip. |  | Yes/No | YES
| 2.1.4 | High Vacuum Suction
| a | Nozzle easily disconnected. |  | Yes/No | YES
| b | Nozzle autoclavable up to 135°C. |  | Yes/No | YES
| c | Degree of rotation for the Nozzle. (degree) |  |  | 360O
| d | Suction intensity can be controlled. |  | Yes/No | YES
| e | Supplied with 2 spare tips (Narrow & Wide nozzle) |  | Yes/No | YES
| f | Light barier switch or button switch for holder. |  | Yes/No | YES
| g | Suction tubing easily disconnected. |  | Yes/No | YES
| h | Continuous suction for not less than 2 hours. |  | Yes/No | YES
| i | Any other features:
2.2 | Doctor’s Instrumentation
| 2.2.1 | Instrumentation
| a | Carried by an adjustable swivel arm with braking system. |  | Yes/No | YES
| b | Connected to the dental chair base moving parallel to the chair. |  | Yes/No | YES
| c | Any other features:
| 2.2.2 | Control Switches
| a | Laminated / Foil Switch. |  | Yes/No | YES
| 2.2.3 | X-Ray Viewer
| a | Attached / Built In |  | Yes/No | YES
| 2.2.4 | Instruments Tubings
| a | Easy disconnected for disinfection. |  | Yes/No | YES
| 2.2.5 | High Torque Turbine
| a | Make
| b | Model
| c | Degree of rotation for the Handpiece. (degree) |  |  | 360O
| d | Handpiece autoclavable up to 135°C. |  | Yes/No | YES
| e | Handpiece non-drip and non-retraction. |  | Yes/No | YES
| f | Handpiece and tubing should be with fibre optic system. (Not less than 25,000 Lux) |  | Yes/No | YES
| g | Handpiece must be push button chuck system. |  | Yes/No | YES
| h | Speed of the handpiece (rpm) |  |  | ≥ 350,000 rpm
| i | Easy Cleaning. Please specify material used. |  | Yes/No | YES
| j | Supplied with three (3) handpieces. |  | Yes/No | YES
|  | (a) Two (2) high torque handpieces. |  | Yes/No | YES
|  | (b) One (1) high torque miniature head turbine handpiece. |  | Yes/No | YES
| k | Any other features:
| 2.2.6 | Micromotor
| a | Make
| b | Model
| c | Type of Micromotor |  |  | electric
| d | Speed (min. 2,000 rpm: max.40,000 rpm) |  |  | ≥ 2000 - 40000 rpm
| e | Micromotor should be with LED fibre optic system. |  | Yes/No | YES
| f | Maintain a high torque at all intensity |  | Yes/No | YES
| g | Motor cannot be demagnetised |  | Yes/No | YES
| h | Clockwise and anticlockwise rotation |  | Yes/No | YES
| i | Internal cooling water spray |  | Yes/No | YES
| j | Must be supplied with the following handpieces (3 units):
|  | (a) Two (2) contra-angle handpieces 1 : 1 with light system. |  | Yes/No | YES
|  | (b) One (1) straight handpiece 1 : 1. |  | Yes/No | YES
| m | All handpieces autoclavable up to 135°C. |  | Yes/No | YES
| n | All contra-angle handpiece fibreoptic system |  | Yes/No | YES
| o | All handpiece integrated water spray. |  | Yes/No | YES
| p | Any other features:
| 2.2.7 | Light Cure Unit
| a | Make
| b | Model
| c | Light Switch located at the handpiece of the unit. |  | Yes/No | YES
| d | Automatic signal sounds can be programmed. |  | Yes/No | YES
| e | Light Intensity must be minimum 1,000 mW/cm2. |  |  | ≥ 1,000 mW/cm2
| f | Light source - LED. |  | Yes/No | YES
| g | Supplied with an optical filter. |  | Yes/No | YES
| 2.2.8 | Ultrasonic Scaler
| a | Make
| b | Model
| c | Magneto-Strictive System / Piezometer System with 28,000 Hz or more. |  | Yes/No | YES
| d | Handpiece or handpiece sleeves autoclavable up to 135°C. |  | Yes/No | YES
| e | 6 detachable working tips for scaling. |  | Yes/No | YES
|  | (a) Buccal / lingual and proximal ( 2 tips ). |  | Yes/No | YES
|  | (b) Sub-gingival ( 4 tips ). |  | Yes/No | YES
| f | Supply with additional one scaler handpiece. |  | Yes/No | YES
| g | To be equipped with light LED fibre optic . |  | Yes/No | YES
| h | Any other features:
| 2.2.9 | Multiple Function Syringe
| a | Make
| b | Model
| c | Function for: | Air | Yes/No | YES
|  |  | Water | Yes/No | YES
|  |  | Spray | Yes/No | YES
| d | Handpiece autoclavable up to 135°C. |  | Yes/No | YES
| e | Degree of rotation for the Handpiece. (degree) |  |  | 360O
| f | Nozzle easily disconnected. |  | Yes/No | YES
| g | Non-drip and non-water retraction. |  | Yes/No | YES
| h | Material of handpiece: |  | Please state | To state by tenderer
| i | Any other features:
2.3 | Cuspidor Unit
a | Attached to the left side of the chair and moves together with the chair. |  |  | Yes/No | YES
b | Consist of spittoon, automatic tumbler filler, saliva ejector and high volume suction. |  |  | Yes/No | YES
c | Electric safety cut-off/limit switch. |  |  | Yes/No | YES
d | Durable spittoon, easy cleaning and cross infection control. |  |  | Yes/No | YES
e | Spittoon bowl one piece type. |  |  | Yes/No | YES
f | Saliva ejector and high volume suction swivelled for easy access by the assistant. |  |  | Yes/No | YES
g | Automatic Control filter activated by foil switches for tumbler. |  |  | Yes/No | YES
h | Fitted with a waste trap. |  |  | Yes/No | YES
i | Any other features:
2.4 | Operating Light
a | Make
b | Model
c | Synchronised movement with chair in vertical direction. |  |  | Yes/No | YES
d | Spring loaded arm with locking brake and height adjustment. |  |  | Yes/No | YES
e | Type of reflector: |  |  |  | Parabolic mirror
f | Light intensity : (lux) |  |  |  | 9,000 - 24,000 lux
g | Halogen bulb or LED of low voltage and easy to replace. |  |  | Yes/No | YES
h | Colour temperature at all intensity : (Kelvin) |  |  |  | 5,500 K
i | Operated by foot control or sensor touch control switch. Please specify |  |  | Yes/No | YES
j | Any other features:
2.5 | Operating Stool
a | Stool for operation : Good lumbar support and adjustable backrest. |  |  | Yes/No | YES
b | Stool for assistant : Good lumbar support and adjustable backrest which can function as armrest and adjustable. |  |  | Yes/No | YES
c | Wide stable base on 5 castors. |  |  | Yes/No | YES
d | Height adjustable: (mm) |  |  |  | 470 - 630 mm
e | Any other features:
2.6 | Independent Water Storage (Self Contained Water System) |  |  |  | YES
a | Make
b | Size
c | Material made
3 | GENERAL
3.1 | Component items of dental unit, like chair, dental unit, operating light and stool are from one manufacturer |  |  | Yes/No | YES
3.2 | Equipment suitable for 230V/50Hz/1Ph |  |  | Yes/No | YES
3.3 | Upholstery colours available:
34 | Utility requirements: (Give details)
| 3.4.1 | Water
| 3.4.2 | Drain
| 3.4.3 | Compressed Air
| 3.4.4 | Suction
| 3.4.5 | Electricity Supply
4 | WARRANTY
4.1 | Two (2) years warranty against manufacturing defects
4.2 | Manufacturer should be able to supply spare parts for 8 to 10 years when the particular unit ceased production
C. | CENTRAL COMPRESSED AIR SYSTEM (DENTAL AIR GRADE)
1 | COMPRESSED AIR SYSTEM
1.1 | Air delivery (litre/min) |  |  |  | ≥ 500 l/m
1.2 | Max. operating pressure (in kPa/bar)
1.3 | Min. operating pressure (in kPa/bar)
1.4 | Automatic operation with manual changeover
1.5 | Standby Unit |  |  | Yes/No | YES
1.6 | Oiless and fitted for dryer |  |  | Yes/No | YES
1.7 | Dental Air Grade |  |  | Yes/No | YES
2 | COMPRESSOR
2.1 | Manufacturer
2.2 | Country of Origin
2.3 | Model (please provide catalogue)
2.4 | Type
2.5 | Noise suppressors for quite operation |  |  | Yes/No | YES
2.6 | Vibration insulation type |  |  | Yes/No | YES
2.7 | Cooling fan for compressor & motor |  |  | Yes/No | YES
2.8 | Air intake filter with silencer |  |  | Yes/No | YES
2.9 | Safety Valve: Brand/Type
2.1 | Shut-off Valve: Brand/Type
2.11 | Non-return Valve: Brand/Type
2.12 | Drain Trap: Brand/Type
2.13 | Pressure Gauge: Brand/Type
2.14 | Testing Flange |  |  | Yes/No | YES
2.15 | Automatic Pressure Switch with Unloaded Start : Type
2.16 | Safety Motor Switch: Type
2.17 | Power Supply: V/Phase/Amp./kW
2.18 | Dimension :W x L x H (mm)
2.19 | Weight (kg)
3 | AIR DRYER SYSTEM
3.1 | Refrigerant dryer or dessicant dryer
3.2 | Capacity
3.3 | Automatic Operation |  |  | Yes/No | YES
3.4 | Refrigerant dryer
Make and Model
Equipped for wall mounting
Capacity of flow
Power supply
Capacity of refrigeration system (kW)
Separator
Automatic Condensate Drain trap
3.5 | Dessicant dryer
Make and Model
After cooler & separator
Desiccant material
Power Supply
Built-in automatic regeneration process for desiccant
Drain
4 | AIR TANK RESERVOIR
4.1 | Capacity of tank (Not less than 20 litres). |  |  |  | ≥ 20 liters
4.2 | Internal lining of corrosion prevention material : Type
4.3 | Working pressure range
4.4 | Dimension (W x L x H) mm
5 | VALVES, FITTINGS, PIPINGS
5.1 | Make/Manufacturer of copper pipe
| Capped and sealed degreased medical grade copper pipe |  |  | Yes/No | YES
5.2 | Make/Manufacturer of copper pipe jointing fitting
| Medical gas fitting, delivered to site individually sealed in bag or box. |  |  | Yes/No | YES
5.3 | Make/Manufacturer of valve
| Material of fittings & valves |  |  | Please state | copper/bronze/brass/stainless steel
6 | WATER FILTER IN JUCTION BOX
6.1 | Make
6.2 | Material
6.3 | Filter type
6.4 | Particle Filteration Size (micron) |  |  |  | 5 micron
7 | CONTROL PANEL FOR COMPRESS AIR SYSTEM
7.1 | Make of cubicle
7.2 | Material
7.3 | Dimension (W x L x H) mm
8 | BACTERIAL FILTER
8.1 | Make
8.2 | Country of Origin
8.3 | Type
8.4 | Particle Removal (mg/m3) |  |  |  | 0.01 mg/m3
9 | AIR INTAKE FILTER
9.1 | Make
9.2 | Country of Origin
9.3 | Type
9.4 | Particle Filteration Size (micron) |  |  |  | 5 micron
10 | DUST FILTER
9.1 | Make
9.2 | Country of Origin
9.3 | Type
9.4 | Particle Filteration Size (micron) |  |  |  | 1 micron
D. | DENTAL SUCTION AIR SYSTEM
1 | SUCTION AIR SYSTEM
1.1 | Capacity (litres/min) |  |  |  | ≥ 1500 l/m
1.2 | Max. operating pressure
1.3 | Min. operating pressure
1.4 | Automatic operation with manual changeover |  |  | Yes/No | YES
1.5 | Standby Unit |  |  | Yes/No | YES
1.6 | Dry Suction system to minimize bacterial contamination in pipeline |  |  | Yes/No | YES
2 | SUCTION
2.1 | Manufacturer
2.2 | Country of Origin
2.3 | Model (please provide catalogue)
2.4 | Type
2.5 | Underpressure limit valve
2.6 | Noise suppressors for quite Operation
2.7 | Dry & oil free
2.8 | Vibration insulation type
2.9 | Electricity supply (V/Phase/Amp./kW)
2.1 | Dimension (W x L x H) mm
2.11 | Weight (kg)
3 | VALVES, FITTING, PIPINGS FOR DENTAL SUCTION AIR SYSTEM
5.1 | Make/Manufacturer of copper pipe
| Capped and sealed degreased medical grade copper pipe |  |  | Yes/No | YES
5.2 | Make/Manufacturer of copper pipe jointing fitting
| Medical gas fitting, delivered to site individually sealed in bag or box. |  |  | Yes/No | YES
5.3 | Make/Manufacturer of valve
| Material of fittings & valves |  |  | Please state | copper/bronze/brass/stainless steel
4 | CONTROL PANEL FOR DENTAL SUCTION AIR SYSTEM
4.1 | Make of cubicle
4.2 | Material
4.3 | Dimensions (W X L X H) mm
5 | SEPARATOR FOR HIGH SUCTION SYSTEM
5.1 | Make and Model |  |  | Please state
5.2 | Fitted with amalgam seperator which comply to EN ISO 11143. (Please provide catalogue for the equipment) |  |  | Yes/No | YES
5.3 | Seperator easily cleaned |  |  | Yes/No | YES
6 | BACTERIAL FILTER
6.1 | Make
6.2 | Country of Origin
6.3 | Type
6.4 | Particle Removal (mg/m3) |  |  |  | 0.01 mg/m3
7 | HEAVY DUTY WATER FILTER
7.1 | Make
7.2 | Material
6.3 | Filter type
6.4 | Particle Filteration Size (micron) |  |  |  | 5 micron
6.5 | Dimension (W x L x H) mm
E. | DENTAL TECHNICIAN WORKING UNIT
1 | Manufacturer
2 | Country of Origin
3 | Model
4 | Dimensions (m)
5 | Material of Work Table : scratch & acid resistant |  |  | Yes/No | YES
6 | Vertically adjustable Fluorescent Light (Rating:lux) |  |  | Yes/No | YES(1800-6000lux)
7 | Posture Guide with arm rest |  |  | Yes/No | YES
8 | Waste Collector Drawer |  |  | Yes/No | YES
9 | Work Tray Holder |  |  | Yes/No | YES
10 | Safety Glass Shield |  |  | Yes/No | YES
11 | Gas Burner with Thermo Electrical Safety |  |  | Yes/No | YES
12 | Dust Extraction System |  |  | Yes/No | YES
13 | Micromotor Speed 2,500-40,000 rpm |  |  | Yes/No | YES
14 | Laboratory Chair |  |  | Yes/No | YES
15 | Utility Requirements: (Give details)
| 15.1 | Gas
| 15.2 | Drain
| 15.3 | Compressed Air
| 15.4 | Electrical Supply
Signature of Sub-Contractor
Name and address (with stamp)
Date
Signature of Witness
Name and address (with stamp)
Date

## 2.Alternative Offers
ALTERNATIVE OFFERS
The tenderer must list hereunder details of any alternative offers.
Alternative offers made after the closing date of tendering will not be considered.
NO. | DESCRIPTION OF ITEM/FEATURES | QUANTITY | DIFFERENCE IN PRICE
Total price difference involved, if any
RM
(Ringgit: |  |  | )
Signature of Nominated Sub-Contractor (with stamp)
Name:
Date: